## Slide 1: D1C/b

## Slide 2: Summary
The Dulcometer D1C/b is the latest in measurement and control from ProMinent. The D1C/b replaces the D1C/a in many applications. Features and exceptions are listed on the following slides.

## Slide 3: Key features – D1Cb
One analog output (record or control)
One digital input
Feed forward signal – pulse only
Trend display (bar graph)
Calibration log (last 10 events)
Error log (last 10 events)
No pH compensation
P/PID control

## Slide 4: Key features – D1Cb
2 pulse frequency outputs
2 power relays (limit, control, timer)
1 alarm relay
Temperature compensation for pH and conductivity
Power supply 90-253 VAC
Same measured variables as D1C/a

## Slide 5: Exceptions
Use D1Ca when:
2 Analog Outputs are needed
pH Compensation is needed (chlorine)
Feed Forward via Analog Signal
Temperature via Analog Signal
Servo motor control

## Slide 6: HACH

## Slide 7: Hach Cl-17 Colorimetric
Top competitor for amperometric
Purchased Great Lakes years ago
Has both colorimetric and amperometric
Most widely used automated equipment for measuring chlorine
Hach is well known for many other testing parameters
Cl-17 is industry standard

## Slide 8: Hach Cl-17 Colorimetric
Uses DPD Chemistry
Only colorimetric and amperometric titrations are accepted methods by EPA
Range is 0-5 ppm of Free or Total Cl
D1C will measure up to 100 ppm of Free
Takes reading every 2.5 minutes

## Slide 9: Hach CL-17 Colorimetric
Analysis is unaffected by pH…..because…..
Free Chlorine – is buffered to 6.3-6.6
Total Chlorine – is buffered to 5.1
So…pH does come into play!!!
Reagents are good for 30 days
Difficult samples need frequent cleaning

## Slide 10: CL-17 Interferences
Bromide
Chlorine Dioxide
Permanganate
Hexavalent Chromium
Dissolved Oxygen
Hydrogen Peroxide
Iron. Copper. Manganese

## Slide 11: CL-17 Interferences
Fingerprints on vials
Dirt and stains
Reagent residue
Dirty photometer lens
Suspended Solids
Bubbles

## Slide 12: Amperometric Interferences
Surfactants
Low conductivity
Dirty Water
Organics
High flow rate past the sensor
Bubbles

## Slide 13: CL-17 Operation
First sample is used as zero point
Compensates for color and turbidity
Indicator and buffer are added
Magnetic stirrer mixes sample and reagents
Colorimeter measures light intensity
Sample is flushed and next sample is taken
  Peristaltic pump
  Colorimeter
  Reagents

## Slide 14: CL-17 Operation
One 4-20 mA output for recording
2 alarm relays
One month supply of reagents
On/Off or proportional control
IP 62 Enclosure

## Slide 15: D1C vs Cl-17
Pricing is comparable to QuickPick
More control parameters with D1C
PID control (D1C)
Two analog outputs (D1C)
Feed forward (D1C)
Correcting value (D1C)
Set up D1C trial next to Cl-17

## Slide 16: Compliance
ProMinent and Hach amperometric technologies are in compliance with EPA Method 334.0
Determination of Residual Chlorine in Drinking Water Using an On-Line Chlorine Analyzer
http://www.epa.gov/ogwdw000/methods/pdfs/methods/met334_0.pdf

## Slide 17: Pricing
Hach CL-17
  $3100 – 3400
CLF10 (Free)
  $3500-4500
CLT10 (Total)
  $3900-4500