# APPENDIX E

Standards Preparation for Sensor Calibration

Recommended Standard solutions:

For pH

Use pH buffer solutions 7 and 10 for two point calibration.  Remove pH sensor from its mount and place in container holding buffer solution.   Rinse sensor and container with distilled water thoroughly before filling with second buffer solution.

For Temperature

A two point calibration can be achieved using an ice bath (32 deg. F) for the low point and a room thermometer for the high point.   The sensor should be removed from its mounting and allowed to stabilize for several minutes before calibration at each value.

For Ammonia

Recommended Standard:   Ricca Chemical Co. Cat. No. 615-32 Ammonia Standard

CAS No. 12125-02-9

1.00 ml = 1.00 mg    NH3 (1000 ppm NH3)

[1.00 ml = 822.4 ug   N (822.4 ppm N)]

NH4Cl in water

Use the ammonium chloride standard solution above and make up standard solutions for two point calibration.   To ensure proper operation and for the initial ammonia sensor calibration, prepare two test solutions that differ in their ammonia concentration by at least a factor of 10.   The standards should be made up to cover the expected operation range.   2 ml of 10 mol caustic soda should be added to each standard prior to calibration to elevate the pH >11 in order to convert any ammonium to ammonia.

Preparation of the two required solutions for ammonia calibration

1.  For stock 100.0 ppm ammonia solution

Pour 10 ml of the recommended NH4Cl (ammonium chloride) standard (1 mg = 1 ml = 1000 ppm NH3) solution into a 100 ml graduated measuring flask partially filled with ammonia free distilled water, mix slightly, and then fill up to the 100 ml mark with ammonia free distilled water.   This is a 100.0 ppm NH3 solution standard which can be used to prepare the actual calibration standards.

2.  For 10-mol caustic soda solution

40g of NaOH (sodium hydroxide) are dissolved in about 80ml of ammonia-free distilled water, in a 100ml measuring flask. The solution will become a little warm. When it has cooled down to 20°C, top up to 100ml with distilled water.

To prepare the ammonia standards and calibrate the sensor

- Close the system sample inlet valve (item 4, Appendix D)

- Rinse the 1000 ml ammonia analyzer calibration column (item 1, Appendix D) with distilled water

- Allow to drain by opening the two valves at the bottom of the column

- Close both bottom valves and fill column approximately 250 ml with ammonia free distilled water.

- Add appropriate amount of the stock 100 ppm ammonia solution through the top of the calibration column based upon the concentration of the standard desired:

  - For   0.1 ppm ammonia add 1 ml

  - For   1.0 ppm ammonia add 10 ml

  - For 10.0 ppm ammonia add 100 ml

- Pipette 2 ml of the 10 mol caustic soda solution through the top of the calibration column.

- Fill the calibration column to the 1000 ml mark with ammonia free distilled water without creating air bubbles.

- Open the valve immediately below the column (valve 2, Appendix D) and allow standard to gravity flow past the ammonia sensor.   Calibrate the sensor per section 2.2 sensor calibration.

- Place valves back in correct position to measure incoming system sample after calibration(s).

NOTE:  Refer to Section 2.2 Sensor Calibration for keypad sequences to perform calibrations.